## A0-003
HSB MRTS HW testlog
Print type | MLA-0387-A0
Serial number | 003
Date | 2018-03-16
Variant
Test location | MK
Test # | Description | Result | Requirement | Pass / fail | Engineer | Date | Comment
1 | Power supply shorts | No | No | Pass | Tla | 2018-03-16
2 | 5V | 4.99 | 4,8 - 5,2 | Pass | Tla | 2018-03-16
3 | 6,5V | 6.57 | 6,2 - 6,8 | Pass | Tla | 2018-03-16
4 | 3,3V | 3.31 | 3,2 - 3,4 | Pass | Tla | 2018-03-16
5 | Vref | 2.5 | 2,49 - 2,51 | Pass | Tla | 2018-03-16
6 | LED | OK | Working | Pass | Tla | 2018-03-16
7 | Display | OK | Working | Pass | Tla | 2018-03-16
8 | Keypad | OK | Working | Pass | Tla | 2018-03-16
9 | Analog input | CH3 ADC 2029 | Vin 1,25V = ADC value 2048 +- 5 | Fail | Tla | 2018-03-16 | Problem is within uC dev board. Opamp circuit is OK.
10 | Analog output | 1.248 | Vout 1,25V +-10mV | Pass | Tla | 2018-03-16
11 | Interlock | OK | Working | Pass | Tla | 2018-03-16
12 | Solenoid | OK | Working | Pass | Tla | 2018-03-16
13 | HV relays | OK | Working | Pass | Tla | 2018-03-16
14 | HV repair | OK | Working | Pass | Tla | 2018-03-16

## A1-003
HSB MRTS HW testlog
Print type | MLA-0387-A1
Serial number | 003
Date | 2018-03-16
Variant
Test location | MK
Test # | Description | Result | Requirement | Pass / fail | Engineer | Date | Comment
1 | Power supply shorts | No | No | Pass | Tla | 2018-03-16
2 | 5V | 4.99 | 4,8 - 5,2 | Pass | Tla | 2018-03-16
3 | 6,5V | 6.62 | 6,2 - 6,8 | Pass | Tla | 2018-03-16
4 | 3,3V | 3.27 | 3,2 - 3,4 | Pass | Tla | 2018-03-16
5 | Vref | 2.498 | 2,49 - 2,51 | Pass | Tla | 2018-03-16
6 | LED | OK | Working | Pass | Tla | 2018-03-16
7 | Display | OK | Working | Pass | Tla | 2018-03-16
8 | Keypad | OK | Working | Pass | Tla | 2018-03-16
9 | Analog input | OK | Vin 1,25V = ADC value 2048 +- 5 | Pass | Tla | 2018-03-16
10 | Analog output | 1,245V | Vout 1,25V +- 10mV | Pass | Tla | 2018-03-16
11 | Interlock | OK | Working | Pass | Tla | 2018-03-16
12 | Solenoid | OK | Working | Pass | Tla | 2018-03-16
13 | HV relays | NA | Working
14 | HV repair | OK | Working | Pass | Tla | 2018-03-16

## A2-003
HSB MRTS HW testlog
Print type | MLA-0387-A2
Serial number | 003
Date | 2018-03-19
Variant
Test location | MK
Test # | Description | Result | Requirement | Pass / fail | Engineer | Date | Comment
1 | Power supply shorts | No | No | Pass | Tla | 2018-03-19
2 | 5V | 4.95 | 4,8 - 5,2 | Pass | Tla | 2018-03-19
3 | 6,5V | 6.59 | 6,2 - 6,8 | Pass | Tla | 2018-03-19
4 | 3,3V | 3.29 | 3,2 - 3,4 | Pass | Tla | 2018-03-19
5 | Vref | 2.497 | 2,49 - 2,51 | Pass | Tla | 2018-03-19
6 | LED | OK | Working | Pass | Tla | 2018-03-19
7 | Display | OK | Working | Pass | Tla | 2018-03-19
8 | Keypad | OK | Working | Pass | Tla | 2018-03-19
9 | Analog input | 2050 | Vin 1,25V = ADC value 2048 +- 5 | Pass | Tla | 2018-03-19
10 | Analog output | 1.249 | Vout 1,25V +- 10mV | Pass | Tla | 2018-03-19
11 | Interlock | OK | Working | Pass | Tla | 2018-03-19
12 | Solenoid | OK | Working | Pass | Tla | 2018-03-19
13 | HV relays | NA | Working
14 | HV repair |  | Working |  | Tla | 2018-03-19